Nephrova Renal Care Pvt. Ltd.

# Biomedical Job Report BME-2026-0311

Equipment: HD-K04 (dialysis machine, Nephrova Kothrud). Engineer: Rahul Pawar. Date: 14 August 2026.

# Complaint

Repeated conductivity alarms on HD-K04 (6 alarms from 11 to 13 August).

# Findings

HD-K04 read a dialysate conductivity of 15.6 mS/cm against 14.0 mS/cm on the reference meter.

The conductivity cell was coated with scale and gave unstable readings.

# Work done

HD-K04 received a conductivity cell replacement followed by calibration against the reference meter.

After repair HD-K04 read 14.0 mS/cm, matching the reference meter.

# Result

Machine released for patient use on 15 August 2026 after a full self test.